2026-08-25 (j.w.) - v2.4.45

ACQUIRE.CPP, START_ACQUIRE

The AutoZero hold on a GC323 is now taken from the Method "AutoZero Time"

field instead of being fixed at 5 seconds. Range 0 to 30 s; a value of 0

keeps the old fixed 5 s. That field is the one previously labelled "Inject

Time", which was unused in this build - it was greyed out in the dialog and

its value was overwritten with a hardcoded 5 s at WAIT_START_INJECT, so

injection timing is not affected. The status line now counts down

"AutoZero: n" in place of "Sampling: n".

ACQUIRE.CPP, EvTimer

The idle oven hold no longer stops heating above 200 C. The duty-cycle loop

called SetOvenTemperature(200) to mean "heater on", but that function is a

setpoint bang-bang, so above about 198 C the call read as "cool down to 200".

It switched the heater off while the caller forced the oven_on flag back to

TRUE, and every heating branch is guarded by if(!oven_on), so the heater was

never re-asserted and the oven coasted down to ambient while the detector

held its temperature. The call now passes t_oven_stored + 10.0. Behaviour

below 200 C is unchanged.

METH_DLG.CPP

The Inject Time field is no longer disabled for MANUAL_INJ builds, and its

range is 0 to 30 s. The four flow-calibration entries, the FlowSetting

message map and ReflectFlowSetting are compiled out: the Method dialog has

had no flow-calibration controls since 2025-08-22 and the leftover

references had been stopping the build.

WPEAK.RC

Field label changed from "Inject Time" to "AutoZero Time"; the Sample Time

and Equil Time labels were narrowed and their edit boxes shifted left to

make room. Range hint and the method printout labels updated to match.

ABOUT.CPP

Version 2.4.45, Build 70.28, copyright year to 2026.

2026-08-16 (j.w.) - ACQUIRE.CPP, START_ACQUIRE

AutoZero at the start of a run is now held for a minimum of 10 seconds.

The hold is still taken from the Method "AutoZero Time" field, but any

value from 1 to 9 s is raised to 10 s; values of 10 s or more are used

as entered. A value of 0 is unchanged - AutoZero drops at the end of the

sample time, as before. Injection timing is not affected.